Novel Flexible Enzyme Laminate-Based Sensor for Analysis of Lactate in Sweat

Eva Tur-García1, Frank Davis1, Stuart D Collyer2, Joanne L. Holmes1, Hugh Barr3 and Séamus P. J. Higson1*

1Cranfield Health, Vincent Building, Cranfield University, Bedfordshire, MK43 0AL, UK.

2Microarray Ltd, Colworth Science Park, Sharnbrook, Bedford, MK44 1LQ, UK.

3Gloucestershire Royal Hospital, Gloucester GL1 3NN, UK

* Corresponding author. Tel: +44(0)-1234-758516. Fax: +44(0)-1234-758380. Email address: s.p.j.higson@cranfield.ac.uk

Abstract

We present a novel amperometric enzyme-based highly flexible biosensor for real-time and non-invasive monitoring of lactate in human sweat for the early detection of pressure ischemia onset. The core of the recognition system is a highly flexible laminate, comprising two highly porous polycarbonate membranes, which provide support for the lactate oxidase enzyme, immobilised via covalent cross-linking. A number of variables were assessed to attempt to optimise the sensors, such as membrane pore size, crosslinking time, crosslinking agent concentration and levels of incorporated enzyme. Oxidation of lactate produces H2O2, which is subsequently determined electrochemically. The transducer comprises a two-electrode system on a single highly flexible polycarbonate membrane, sputter-coated with gold and platinum to render it conductive. The sensor exhibits lactate selectivity with a working range of 0-70 mM, thus covering physiologically relevant concentrations for pressure ischemia and has been shown to be suitable for determination of lactate in PBS, synthetic sweat and human sweat.

Keywords: lactate; pressure ischemia; sweat

1. Introduction

Pressure ischemia is the condition in which skin and underlying tissue necrosis occurs due to malnutrition of the tissues in body areas exposed to continued pressure, or pressure in combination with shear and/or friction (Bansal et al., 2005). People more susceptible to pressure ulcers include those with mobility problems such as bedridden, wheelchair-bound, debilitated or paralyzed patients. People suffering from cardiovascular and neurological diseases can in many cases also suffer from pressure ischemia, since these diseases lead to an impairment in circulation (Bansal et al., 2005).

When the skin undergoes prolonged pressure, the underlying blood vessels become occluded, either partially or totally. As a result of this, oxygen and other nutrients carried in the blood are not delivered in sufficient quantities to satisfy the metabolic demands of the affected tissue. Cells are then obliged to use their own stores of energy through an anaerobic metabolic pathway in order to survive. Consequently, breakdown products from metabolism start to accumulate within both the affected cells and the interstitial spaces. As the levels of energy stores diminish, cellular processes start to fail, the ionic flow across cellular membranes begins to fall and then cell necrosis occurs, with the subsequent formation of a pressure ulcer (pressure sore). Hence, pressure ulcers arise from prolonged tissue ischemia (Polliack et al., 1993).

It is estimated that, in general terms, approximately 9% of hospitalised patients develop pressure sores (Bauer et al., 2006). Moreover, it is well understood that pressure ischemia is related to a notable increase in mortality rates (Lyder and Ayello, 2008). In acute care hospitals, the prevalence of pressure ulcers is approximately 10% (Phillips and Buttery, 2009). It is estimated that the treatment of pressure ulcers costs the UK National Health Service 4% of its budget, a cost of between £1.4 and £2.1 billion per year (Bennett et al., 2004). It is for this reason that pressure ischemia has significant implications in cost and quality of service for the health sector.

One of the most common clinical practices for the early detection of pressure ischemia was based on a periodical inspection of the skin colour of the patient. Later diagnosis methods were based on an alternative approach using tissues as the source of information. The first technique consisted of the monitoring of blood flow, which constituted a slight improvement from the observation of skin colour (Ferguson-Pell and Hagisawa, 1988). It was known that even in early stages of pressure ulcer formation, the tissues responded with an accumulation of metabolites, a decrease in pH and oxygen partial pressure (PO2, which is the amount of oxygen that will bind haemoglobin within the red blood cells) and an increase in partial pressure of carbon dioxide (PCO2) (Matsen et al., 1981). However, the techniques derived from this idea implied the use of invasive methods on areas of risk for the collection of samples. Pressure ischemia is considered a preventable condition that arises from poor clinical management, reduced nursing staff and a lack of warning indicators (Bours et al., 2001). However, in reality it is understood that it is not that easy to prevent because of the many factors that cause it.

Pressure ischemia causes a decrease in oxygen and glucose levels, forcing the cells within the vascularised tissues to change to anaerobic metabolism. Anearobic metabolism of the cells’ glucose reserves leads to the production of more lactic acid, which starts accumulating (Stekelenburg et al., 2008). However, if lactate concentration exceeds or glucose levels fall below a certain threshold level, cell death will occur (Gawlitta et al., 2007; Stekelenburg et al., 2008).

Lactate levels constitute a useful warning indicator for the formation of pressure ulcers. In susceptible patients, pressure ulcers can develop in just one or two hours (Edlich et al., 2004), even though in some cases the damage will only become evident a few days later. Therefore, an immediate and continuous monitoring of lactate is a key factor for in-time intervention by physicians (Smith et al., 2001). This monitoring requires biosensors with direct, simple and highly specific lactate measurements, low response time, inexpensive and with minimal or no sample preparation. Commercially available biosensors for blood lactate have been developed but of course require an invasive method of sample collection. Implantable vascular sensors for in-situ monitoring of lactate levels have been developed, however, these entail a high risk of thrombosis and embolism (Spehar-Deleze et al., 2012). This issue and the current trend for non-invasive diagnostic and real-time monitoring methods have led researchers to search for other body fluids on which to perform lactate measurements and among these, sweat shows the most promise, since it is the most accessible specimen to collect, its collection is non-invasive and its real-time analysis offers valuable physiological information (Coyle et al., 2010). Also, sweat analysis allows specific body areas to be studied (i.e. pressure points).

It was suggested that sweat lactate would be a good marker for evaluating the severity of peripheral occlusive arterial disease which implies a local occlusion of blood supply (Fellmann et al., 1989). Subsequent studies agreed with this and stated that sweat lactate could be a sensitive indicator of damage in soft tissues (Polliack et al., 1993; Polliack et al., 1997; Knight et al., 2001).

In order to achieve continuous monitoring, wearable sensors have attracted increasing interest from the research community and the industry during the last decade (Morris et al., 2009). To date, most of the existing wearable sensors have been developed for the monitoring of physiological parameters of the patient, such as breathing rate, heart rate or temperature (Coyle et al., 2007). Wearable chemo-sensors, although at an early stage, have the potential to measure many more variables related to the wearer’s health. Moreover, the potential of these sensors to perform chemical and continuous measurements of body fluids (i.e. sweat) opens a new insight for medical science by providing valuable real-time feedback information about the patient’s health status, which is the key for preventative healthcare and early diagnosis of conditions such as pressure ischemia.

A review of lactate biosensors is beyond the scope of this paper, however previous workers (Nikolaus and Strehlitz, 2008) have analysed the different aspects within the preparation of amperometric lactate biosensors such as biorecognition elements, methods of immobilisation, mediators and cofactors, and fields in which these sensors can be applied. Most of the sensors utilise the technology developed for glucose monitoring, with lactate enzymes such as lactate oxidase forming the active biorecognition element. Our present work describes the development of a novel flexible biosensor for real-time, continuous and non-invasive sensing of lactate in human sweat. This potentially will lead to development of an external sensor that will be directly applied to the skin of patients as a real-time, non-invasive test for the onset of pressure ischemia.

2. Materials and methods

2.1 Reagents and materials

Ferrocene carboxylic acid, lactate oxidase (from Pediococcus sp., lyophilised powder, ≥20 units/mg solid), sodium L-lactate and bovine serum albumin (BSA, lyophilised powder) were purchased from Sigma-Aldrich (Dorset, U.K.). Sodium dihydrogen orthophosphate 1-hydrate (NaH2PO4•H2O), disodium hydrogen orthophosphate 12-hydrate (Na2HPO4•12H2O), sodium chloride, urea, ascorbic acid, glacial acetic acid, hydrochloric acid and ≈50% glutaraldehyde solution were purchased from VWR BDH (Poole, U.K.). All reagents were used as received unless otherwise stated.

Deionised water for preparation of all solutions was obtained using Purelab® UHQ from ELGA (Marlow, U.K) and had a resistivity of >18 MΩ cm-1. A phosphate buffered saline (PBS) solution was prepared at pH 7.4 containing 5.28 x 10-2 M Na2HPO4•12H2O, 1.3 x 10-2 M NaH2PO4•H2O and 5.1 x 10-3 M NaCl. BSA solutions were prepared at a concentration of 0.1 g mL-1 in PBS. A lactate oxidase/BSA solution was prepared using LOD at the desired concentration dissolved in 0.1 g mL-1 of BSA solution for use within an enzyme laminate. Enzyme/BSA solutions were stored in Eppendorf tubes at -20°C until use.

A suggested chemical formulation of artificial human sweat (Stefaniak and Harvey 2006; Harvey et al. 2010) was made by dissolving 1.55 g L-1 sodium chloride, 0.6 g L-1 urea, 7.81 x 10-3 g L-1 acetic acid and 1.76 x 10-3 g L-1 ascorbic acid in PBS pH 7.4 and a range of lactate concentrations, from 0 to 70 mM, were prepared from this basic solution and analysed with the AuPt electrode system.

Human sweat was collected from a number of healthy volunteers, after running at speeds of 11-13 km h-1 on a treadmill for up to one hour. Samples were collected from various parts of the body using sterile absorptive pads strapped to the body with surgical tape which were removed and then centrifuged to extract the sweat. The pads were then placed in centrifuge tubes and spun to extract the sweat from the pad. Samples were diluted 1:8 in PBS before measurement and tested as soon as possible after collection.

Polycarbonate membrane disks of pore sizes from 0.015 to 1.0 µm diameter with a thickness of 10 µm were purchased from both Whatman (Maidstone, U.K.) and Millipore (Consett, U.K.). Glass microscope slides (ground edges, twin frosted glass 76 mm x 26 mm, 0.8 mm to 1.0 mm thick) were purchased from Fisher Scientific (Loughborough, U.K.). Araldite® epoxy resin (Rapid setting), multicore wires and silver conductive paint were purchased from RS Components (Corby, UK).

2.2 Equipment

A Uniscan (Buxton, Derbyshire, U.K.) PG580 with dedicated PC-driven software was used for electrochemical studies such as linear and cyclic voltammetry as well as chronoamperometry experiments. Sputter coating of metals onto the polycarbonate membranes was performed by using an AGAR B7341 Automatic Sputter Coater from Agar Scientific Ltd. (Essex, U.K.) in conjunction with an external Pfeiffer Rotary Vane Pump from Pfeiffer Vacuum Ltd. (Newport Pagnell, U.K.).

2.3 Manufacture of electrodes

A 25 mm polycarbonate disk covered by a plastic film template carefully cut to cover the area for the working electrode and the non-coated region between electrodes (following the pattern shown in figure 1a) was sputter-coated with gold (120 seconds exposure) to deposit the counter electrode. The tape was then carefully removed and the resulting sputter-coated membrane covered by a plastic film template carefully cut to cover the polycarbonate disk but leave the working electrode area exposed. The template with the membrane beneath was fixed to the glass plate of the sputter-coater with electrical tape. The disk was then sputter-coated with platinum (150 seconds exposure) and then the plastic template was removed, leaving a structure comprised of two electrodes, an inner platinum-coated working electrode with an outer gold counter/reference electrode on a single membrane separated and isolated from each other by a region of non-coated membrane (figure 1a). The times for deposition were determined by depositing the metals onto the surface of polycarbonate membranes and then measuring conductivities of the resulting films, conductivities increased with deposition time until 120 seconds (gold) or 150 seconds (platinum), with further deposition leading to only minimal increases in conductivity.

Two lengths of wire of 10-15 cm were cut and stripped of the insulating plastic coat for 3 mm on each end. These were then attached to the metallised area of both working and counter/reference electrodes (as shown in figure 1b) using silver conductive paint to make a connection between the wire and the membrane. The silver conductive paint was allowed to set for one hour and then the wires were fixed to the membrane surface by applying epoxy resin, allowing it so set until dry, to ensure they were insulated from solution and to add structural integrity. Finally the electrode was attached to a glass slide or a flexible plastic sheet using epoxy, not only to improve integrity but also to minimise any flow of lactate through the lower membrane to the laminate.

An enzyme-crosslinking mix was made by rapidly mixing 6 µl of the enzyme/BSA solution mas with 3 µl of 5% or 10% glutaraldehyde (PBS). The enzyme laminate was prepared with two 5 x 9 mm polycarbonate membranes and the crosslinking solution placed between them. This was then lightly compressed between two microscope slides for five minutes and then rinsed in PBS to remove unreacted glutaraldehyde and stop the reaction. Once a robust enzyme laminate was obtained, it was placed on top of the working electrode (figure 1b). The membranes of the enzyme laminate were cut so they would mimic the shape of the working electrode without making contact with the counter/reference electrode area. Finally, the enzyme laminate was attached to the electrode by placing thin strips of insulating tape along the edges without encroaching on the counter/reference electrode area. Figure 1c shows a schematic of the final arrangement along with a photograph (figure 1d) of the electrode after attachment of the wires.

2.4 Chronoamperometric analysis

The electrode system was connected to the potentiostat with the inputs for the counter and reference electrode being merged and then both electrodes (working and counter/reference) were connected to their corresponding inputs in the potentiostat. The sensor was placed in a 25 mL beaker and 15 mL of PBS were initially added. The sensor was then set to polarise at +650 mV until a baseline was reached. The working electrode was polarised at +650 mV vs. and the resulting current was monitored. This potential of +650 mV was chosen since it is known to be the optimum value for the oxidation of hydrogen peroxide (Updike and Hicks 1967). Initially a high current followed by a dramatic decay in the response is observed due to the polarisation and relaxation of the electrical double layer. Once this was complete, the current reached a constant baseline in terms of steady state current (constant current against time). The time for the polarisation to conclude varied depending on the nature of the experiment (i.e. the hydration status of the enzyme laminate). Once a stable baseline being obtained, the system was ready to use for analysing different samples. 15 mL of each lactate solution (in PBS only, in synthetic sweat or diluted human sweat) were sequentially added for their analysis. After the experiment, the current/time profile was analysed for each substrate concentration. In order to calculate normalised current responses, the baseline current was subtracted.

3. Results and discussion.

3.1 Optimisation of the electrodes and laminate

Studies were performed focused on determining the reproducibility of the electrode-system response so as to assess the quality of the hand-manufactured electrode system in terms of providing a reproducible electrochemical response between different sensing devices. The electrodes were assessed using cyclic voltammetry between -0.4 and +0.8 (vs. an Ag/AgCl reference electrode) at a sweep rate of 20 mV s-1 (Gornall et al., 2009). Each solution was tested on five different sensors in order to obtain an average current response. The voltammograms obtained from the analysis of each sensor for the range of FA concentrations are shown in figure S1 (Supplementary information) as well as the average positive (anodic) and negative (cathodic) peak currents for each concentration figure S2 (Supplementary information). RSD values of 5.4-9.8% indicated that the variation in electrochemical properties between different hand-manufactured electrode systems was relatively small and, therefore, would take a minimum part in the overall variability within the current response of the system.

Initial studies were made on the effect of variation of membrane pore size on current density. As can be seen (figure 2) the initial experiments using 1 µm pore size membranes gave a good current density, indicating there is throughput of lactate, oxygen and hydrogen peroxide. However as can be seen, the increase response falls off rapidly above 0.4 mM lactate, which is an issue since the clinical ranges of lactate that are required to be measured are much higher than this. One obvious explanation is that the enzyme is reaching saturation (Km of lactate oxidase is approximately 0.7 mM) and that should diffusional barriers to the substrates (lactate and oxygen) be put in place, as has been used for many other sensing systems, then the linear range could be increased. Therefore less porous membranes were utilised as the upper membrane (furthest from the electrode) in these laminates. Membrane laminates were constructed using three different pore sizes, applied to the working electrodes and tested in lactate solutions. Smaller pore membranes, although reducing current response and sensitivity appear to increase the linear range and therefore the smallest pore size (0.015 µm) shows potential for increasing the linear range especially if other variables are optimised. Therefore, unless otherwise stated, 0.015 µm pore membranes were used as the upper membrane.

In an attempt to determine whether a platinum working electrode is required, a similar system containing gold working and counter/reference electrodes was constructed since these could be deposited in a single sputter coating step. However when an enzyme laminate (1 µm polycarbonate membrane) was applied to this membrane and the current response to lactate measured, current densities of approximately a third of those obtained for platinum working electrodes were obtained, confirming the catalytic properties of platinum towards hydrogen peroxide.

These initial results were obtained using an enzyme laminate containing 830 U mL-1 of lactate oxidase and crosslinked using a 5% glutaraldehyde solution with an incubation time of 8 minutes, this being the shortest time that gave a firm, physically stable laminate. It has been previously reported that the degree of cross-linking increases with the incubation time, during which the reaction from glutaraldehyde with amino groups (mainly from lysine residues) from lactate oxidase and BSA continues to take place (House et al., 2007). Cross-linking time, like many other parameters, passes through an optimum value. However, this incubation time varies depending on many factors, such as the chosen immobilisation method, materials or substrates employed and their concentration, physical characteristics and the concentration of the enzyme to be immobilised as well as other experiment conditions such as pH or temperature.

Enzyme laminates were allowed to crosslink for 8-30 minutes; these were then applied to the platinum working electrodes and the response to lactate measured. Both upper and lower membranes in the laminate were of 1 µm pore radii. This pore size was chosen in order to avoid any lactate diffusion restrictions due to the membrane employed and to make the results from the electrochemical characterisation of the laminate depend exclusively on the properties of the cross-linked matrix. Figure 3 shows data obtained for three incubation times; increasing the incubation from 8 minutes to 12 minutes increased sensor response, however a further increase lowered it. Longer times led to even lower responses and after 30 minutes the resultant sensor displayed minimal response to lactate. Therefore 12 minutes was chosen as the optimal incubation time.

The amount of cross-linker and the glutaraldehyde to LOD/BSA ratio also has an optimum point as it has an important influence on the enzyme activity and operational stability in a similar way to the incubation time (Sheldon, 2011), with low levels leading to insufficient crosslinking and enzyme leaching, while higher levels rendered the laminates too rigid and losing their activity (Matijošytė et al., 2010). Therefore, the optimum amount of cross-linker is reached when a sufficiently robust matrix is created with enough cavities to avoid excessive mass transfer limitations, while the enzyme preserves its activity and stability. However, this optimum value of cross-linker will depend on the enzyme employed, since the number of free amino groups (mainly from lysine residues) on their surface available to interact with the cross-linker varies from enzyme to enzyme (Sheldon, 2011). Therefore the concentration of glutaraldehyde was increased to 10% (higher concentrations were tried but the resultant reaction was too quick and the mixture solidified before it could be applied as a laminate).

Figure 4a shows a clear increase in the linear range of the sensor when a range of concentrations of lactate was applied. However there is a noticeable drop in sensitivity compared to 5% glutaraldehyde crosslinked solutions. This is due to the more crosslinked laminate restricting lactate flow and potentially also due to glutaraldehyde reacting with lysine units in the enzyme and reducing its activity. It appears that increasing the concentration of glutaraldehyde has the largest single effect on improving the linearity of the system and therefore 10% glutaraldehyde was used in further experiments.

One challenge to the use of the higher concentration of glutaraldehyde is the reduced sensitivity, however other workers have reported that that an increase in the amount of immobilised enzyme also improves the signal height and the stability of the sensor (Gorton, 2005). This is due to the fact that for relatively high substrate concentrations, the enzyme turnover rate (enzyme kinetics) is the limiting factor of the signal produced. However, if a sufficient enzyme loading is achieved, substrate molecules reaching the enzyme layer within the laminate will be rapidly converted and, then, the signal produced by the sensor is limited only by the diffusion rate of the substrate to the enzyme layer. At this point, the enzyme activity does not affect the signal and, thus, this activity can decrease through denaturation or blocking of the active site without it affecting the signal height, thus keeping a high apparent stability (Gorton, 2005). Therefore, the amount of enzyme in the laminate precursor solution was doubled and as figure 4b shows, this led to a large increase in signal whilst still retaining a working range out to 70 mM lactate. It appears therefore the most effective method for laminate synthesis is to utilise higher levels of crosslinking reagent and enzyme.

3.2 Measurements in synthetic sweat

In order to assess the effect of possible interferents such as ascorbate, a synthetic sweat solution containing a typical physiological concentration of ascorbate (Stefaniak and Harvey 2006; Harvey et al. 2010) amongst other materials, formulated as described earlier was used to make up lactate solutions. Enzyme laminate electrodes were formulated using optimised conditions (10% glutaraldehyde, 12 minutes crosslinking time and 1667 U mL-1 of lactate oxidase) and used to measure lactate. Figure 5a shows current time responses for an electrode in alternating solutions of 10 mM lactate in PBS and synthetic sweat. As can be seen there is minimal interference from the matrix, with the peak currents being within 2% of each other. A calibration profile for lactate in synthetic sweat was also obtained and shows very similar behaviour to that obtained in PBS.

3.3. Assessment of human sweat lactate levels.

Human sweat samples were diluted 1:8 in PBS (so that there would not be large variations in ionic strength) and these were assessed using lactate electrodes. Two separate electrodes were used and profiles are shown in figure 6a. Although the two electrodes gave differing responses since they were hand fabricated, it can be shown their responses are internally consistent by first calibrating them with solutions of lactate in PBS. After this step, two separate dilutions of sweat in PBS were applied to the electrodes and again consistent current/time profiles obtained.

Figure 6b shows similar experiments made using sweat collected from differing parts of the body. As can be seen sweat samples taken from the chest displayed higher lactate levels (15 mM once dilution is taken into account) than obtained from the neck or back (10 mM). It has already been reported by a number of other workers that sweat lactate can vary depending on the site of collection as reviewed earlier (Derbyshire et al 2012).

Figure 6c shows the effect of increasing the exercise levels on lactate. Running for one hour produced a reading of more than double that obtained after 30 minutes (28 mM and 13 mM respectively). Whether this is a measure of body lactate levels or just a higher metabolic activity of the sweat gland is still being debated as previously reviewed (Derbyshire et al 2012). However intense exercise is known to produce hypoxia, thereby increasing anaerobic metabolism and the production of lactic acid.

Conclusions

A flexible sensor based on a platinum working electrode and gold counter/reference electrode on polycarbonate membrane has been developed. Different parameters involved in the construction of the enzyme laminate were optimised in order to increase the reproducibility and linearity of the response. A combination of high levels of enzyme and crosslinking agent and a 12 minute incubation time were found to give the best results with workable ranges between 0-70 mM lactate and good sensitivity. The developed sensing design can selectively detect lactate at physiologically relevant concentrations in PBS or synthetic sweat with no significant interference observed. Diluted human sweat samples were analysed in the sensing system, which provided useful information on sweat lactate variability with body location and exercise levels.

Further work will focus on analysing undiluted sweat obtained first from healthy volunteers and later for hospital patients. Other work will analyse the potential for improving the electrode performance by incorporation of mediators in the system. Finally attempts will be made to use the electrodes in continuous monitoring of sweat lactate.

Acknowledgement

The authors would like to thanks Gloucester NHS Trust for funding for ELTG.

References

Bansal, C., Scott, R., Stewart, D. and Cockerell, C. J. 2005, Int. J. Dermatol., 44, 805-810.

Bauer, J. D., Mancoll, J. S. and Phillips, L. G. 2006, "Pressure sores", in Thorne, C. H., Beasley, R. W., Aston, S. J., et al eds. Grabb and Smith's Plastic Surgery, Sixth ed, Lippincott Williams & Wilkins, Philadelphia, pp. 722-729.

Bennett, G., Dealey, C. and Posnett, J. 2004, Age and Ageing, 33, 230-235.

Bours, G., Laat, E., Halfens, R. and Lubbers, M. 2001, Int. Care Med., 27, 1599-1605.

Coyle, S., Wu, Y., Lau, K. -., Brady, S., Wallace, G. and Diamond, D. 2007, "Bio-sensing textiles -wearable chemical biosensors for health monitoring", IFMBE Proceedings, 13, 35-39.

Coyle, S., Benito-Lopez, F., Byrne, R. and Diamond, D., 2010, On-body chemical sensors for monitoring sweat. In Wearable and Autonomous Biomedical Devices and Systems for smart Environment. Ed. A. Lay-Ekuakille, Springer.

Derbyshire, P. J. Barr, H. Davis, F. Higson, S. P. J. 2012. J. Physiol. Sci., 62, 429-440.

Edlich, R. F., Winters, K. L., Woodard, C. R., Buschbacher, R. M., Long, W. B., Gebhart, J. H. and Ma, E. K. 2004, J. Long Term Eff. Med. Implants., 14, 285-304.

Fellmann, N., Fabry, R. and Coudert, J. 1989, Am J Physiol. 257, H395-H398.

Ferguson-Pell, M. and Hagisawa, S. 1988, J. Rehab. Res. Develop., 25, 57-62.

Gawlitta, D., Oomens, C. W. J., Bader, D. L., Baaijens, F. P. T. and Bouten, C. V. C. 2007, " J. Appl. Physiol., 103, 464-473.

Gornall, D. D., Collyer, S. D. and Higson, S. P. J. 2009, Sen. Actuat. B: 141, 581-591.

Gorton, L. 2005, Biosensors and modern biospecific analytical techniques, Elsevier, Amsterdam; Boston.

Harvey, C. J. LeBouf, R. F.; and Stefaniak, A. B. 2010. Toxicol. In Vitro. 24, 1790-1796.

Knight, S. L., Taylor, R. P., Polliack, A. A. and Bader, D. L. 2001, J. Appl. Physiol., 90, 2231-2237.

Lyder, C. H. and Ayello, E. A. 2008, "Pressure ulcers: a patient safety issue.", in Hughes, R. G. ed. Patient safety and quality. An evidence-based handbook for nurses. AHRQ, U.S.

Matsen, F. A.,3rd, Wyss, C. R., King, R. V., Barnes, D. and Simmons, C. W. 1981, Clin. Orthopaed. Rel. Res. 155, 224-230.

Matijošytė, I., Arends, I. W. C. E., de Vries, S. and Sheldon, R. A. 2010, J. Mol. Cat B, 62, 142-148.

Morris, D., Coyle, S., Wu, Y., Lau, K. T., Wallace, G. and Diamond, D. 2009, Sens. Actuat., B: 139, 231-236.

Nikolaus, N. and Strehlitz, B. 2008, " Microchim. Acta, 160, 15-55.

Phillips, L. and Buttery, J. 2009, Nursing Times, 105, 34-36.

Polliack, A., Taylor, R. and Bader, D. 1993, J. Rehab. Res. Develop., 30, 250-259.

Polliack, A., Taylor, R. and Bader, D. 1997, J. Rehab. Res. Develop., 34, 303-308.

heldon, R. A. 2011, App. Microbio. Biotech., 92, 467-477.

Smith, I., Kumar, P., Molloy, S., Rhodes, A., Newman, P. J., Grounds, R. M. and Bennett, E. D. 2001, Int. Care Med, 27, 74-83.

Spehar-Deleze, A., Anastasova, S., Popplewell, J. and Vadgama, P. 2012, "Extreme Physiological State: Development of Tissue Lactate Sensor", Ninth International Conference on Wearable and Implantable Body Sensor Networks BSN, 17-21.

Stefaniak, A. B. and Harvey, C. J. 2006. Toxicol. In Vitro 20, 1265-1283.

Stekelenburg, A., Gawlitta, D., Bader, D. L. and Oomens, C. W. 2008, Arch. Phys. Med. Rehab., 89, 1410-1413.

Updike S. and Hicks G. 1967. Nature, 214, 986-988.

List of figures.

Figure 1. (a) schematic of the Au/Pt electrode (b) application of enzyme laminate (c) schematic and (d) photograph of final sensor layout.

Figure 2. Effect of membrane pore size – current vs. lactate concentration with laminates constructed with upper-membrane porosities of 1, 0.1 and 0.015 µm (average for three of each laminate).

Figure 3. Effect of cross-linking time – current vs. lactate concentration for three different incubation times. Each point is the average of three separate electrodes.

Figure 4. Current vs. lactate concentration for laminates made using 10% glutaraldehyde and two different concentrations of LOD. Each point is the average of four separate electrodes.

Figure 5. (a) Current response over alternating applications of 10 mM lactate in PBS only (PBS) and synthetic sweat (SS). The arrows indicate the moment at which the PBS solution was extracted from the beaker and the lactate test solutions were introduced. (b) Calibration plot for lactate in synthetic sweat. Each point is the average of three separate electrodes.

Figure 6. (a) Current response for two electrodes over applications of lactate in PBS only (PBS) followed by two separate applications of a single human sweat sample (30 mins of running, 1:8 dilution in PBS). (b) current response over applications of lactate in PBS only (PBS) followed by applications of human sweat samples collected from different parts of the body (30 mins of running, 1:8 dilution in PBS). (c) Current response over applications of lactate in PBS only (PBS) followed by applications of human sweat samples collected after two different exercise times (1:8 dilution in PBS).

Figure 1.

Figure 2.

Figure 3.

Figure 4.

Figure 5.

Figure 6.